The ATLAS  SCT End-Cap Detector Modules

| Prepared by : |  |  | Checked by : |  |  | Approved by : |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Z. Dolezal, L. Feld, D. Ferrere, D. Macina, H.-G. Moser, V. O’Shea, T. Smith, S. Snow |  |  |  |  |  |  |  |
|  | for information, you can contact : | L. Feld |  | Tel. +49 761 203 5753 | Fax. +49 761 203 5931 |  | E-Mail Lutz.Feld@cern.ch |

- Distribution:

| Issue Revision | Date | Sheet | Description of Change | Release |
| --- | --- | --- | --- | --- |
| Draft 1 | 8-2-2001 |  | This is the first draft version of this document |  |

1	Technical Specification of the Modules	6

1.1	Electrical	6

1.2	Mechanical	6

1.3	Thermal	6

1.4	Interfaces	6

1.4.1	Electrical	6

1.4.2	Mechanical including Alignment and Cooling	6

1.5	Environment	6

2	Technical Specification of the Components	6

2.1	Detectors	6

2.2	ASICs	6

2.2.1	ABCD	6

2.2.2	VDC and DORIC	6

2.3	Hybrid	6

2.3.1	Flexible Circuit	6

2.3.2	Substrate	6

2.4	Spine	6

2.5	Fan-Ins	6

2.6	Location Washers	6

2.7	Glues	6

3	Assembly Procedures	6

3.1	Procedures	6

3.2	Equipment	6

3.3	Facilities at each Production Site	7

3.4	Qualification of each Assembly Site	7

3.5	Accounting	7

3.6	Packaging and Shipments	7

3.7	Assembly of Modules onto Disks	7

4	Quality Assurance	8

4.1	Components	8

4.2	Hybrids (preliminary)	8

4.2.1	Tests done at vendor/procurement institute	8

4.2.2	Electrical Tests of Hybrids after ASIC Attachment	8

4.2.3	Long Term (Burn-in) Test of Hybrids with ASICs	9

4.2.4	Final Electrical Test of Hybrids	9

4.3	QA during module assembly	9

4.4	QA of the completed module	9

4.4.1	IV curve	9

4.4.2	Thermal cycling	9

4.4.3	Metrology survey	10

4.4.4	IV curve of the completed module	10

4.4.5	Long-term leakage current stability	10

4.4.6	Module electrical confirmation test	10

4.4.7	Long-term electrical stability	10

4.4.8	Final electrical characterization tests	10

4.5	Sampling QA on completed modules	10

4.5.1	Laser Tests	11

4.5.2	Irradiation Tests	11

4.5.3	Beam Tests	11

4.5.4	Source Tests	11

4.6	Definition of Electrical Hybrid and Module Tests	11

4.6.1	Power on Tests / Verification of Response to Hard Reset	12

4.6.2	Clock and Command Reception / Addressing Error Test	12

4.6.3	Bypass Functionality Test	13

4.6.4	Pipeline Efficiency Test	13

4.6.5	Strobe Delay Scan	13

4.6.6	Three Point Estimation of Gain, Noise and Offset	13

4.6.7	Trim Range Scan	14

4.6.8	Response Curve	14

4.6.9	Noise Occupancy Scan	14

4.6.10	Timewalk Scan	14

5	16

6	Project Schedule	16

6.1	End-Cap Construction Schedule	16

6.2	Module Production Responsibilities	16

6.3	Component Supplies	16

6.4	Timetable	16

7	Prototype Module Results	16

7.1	Pre-irradiation at Individual Module Level	16

7.1.1	Mechanical and Thermal	16

7.1.2	Electrical	16

7.2	Pre-irradiation at Combined Modules Level	16

7.2.1	System Test	16

7.3	Post-irradiation	16

7.3.1	Mechanical Integrity	16

7.3.2	Electrical Performance	16

8	References	16

# Technical Specification of the Modules

## Electrical

## Mechanical

## Thermal

## Interfaces

### Electrical

### Mechanical including Alignment and Cooling

## Environment

# Technical Specification of the Components

## Detectors

## ASICs

### ABCD

### VDC and DORIC

## Hybrid

### Flexible Circuit

### Substrate

## Spine

## Fan-Ins

## Location Washers

## Glues

# Assembly Procedures

## Procedures

## Equipment

## Facilities at each Production Site

## Qualification of each Assembly Site

## Accounting

## Packaging and Shipments

## Assembly of Modules onto Disks

# Quality Assurance

This chapter describes the QA procedure to be applied to end-cap modules during and after the construction process. Post-assembly testing validates the quality of the completed module: most tests will be performed on all produced modules, but some destructive or very time-consuming tests are performed on a sample subset of modules.

The procedures and tests described here aim to investigate a wide range of potential failures and problems.  In the later parts of the production, where the production rate is higher and experience of previously produced modules has been gained, the procedures will be revisited and reduced in areas that do not show failures, or, on the other hand, extended in problematic areas.

The results of all tests will be logged to the SCT database. In some cases a simple yes/no result is obtained, for most tests characteristic performance data will be logged in addition to the flag of whether the module passes the relevant QA criteria.

The aim of the QA procedures is to ensure that each SCT End-Cap Module fulfils all aspects of the specification.  The same QA procedures and criteria will be applied at all the end-cap module production sites. The criteria for accepting a module as ‘good’ are for initial use, and will evolve with experience.

## Components

The QA procedures for most of the individual components are dealt with in other documents.  Tests are made during assembly to check that the components have not been damaged, for example in transport or in the first stages of module construction. Principles of these checks are described in another document (Assembly Criteria).  In this chapter, QA procedures of hybrids and modules are described.

## Hybrids (preliminary)

### Tests done at vendor/procurement institute

Hybrid production will comply with industry standards of quality assurance. Furthermore, several tests will be performed to verify resulting quality of the flex, electrical connections, substrate, etc. After population with passive components hybrids will be thermally cycled and afterwards continuity of traces will be tested.

After ASIC attachment basic electrical tests (Confirmation Sequence – see section ???)  are performed to check all wire bonds and basic electrical functionality of ASICs.

### Electrical Tests of Hybrids after ASIC Attachment

Hybrids with ASICs attached will be distributed to 3 hybrid-testing sites (one per cluster). After transport a visual check is made of each hybrid to check that no damage has occurred in transport. For similar reasons basic electrical functionality is tested using the Confirmation Sequence.

### Long Term (Burn-in) Test of Hybrids with ASICs

In order to catch infant mortality problems in the front-end ASICs a longer duration burn-in test is performed on assembled hybrids.

The test consists of a long duration run at elevated temperature. The temperature and length of the run will be adjusted during production in the light of experience gained, but initially a 100-hour test will be made with a temperature measured by the hybrid thermistors of ?? degrees C.  During the test the hybrids are powered, clocked and configured and triggered at the nominal L1A trigger frequency of 100kHz.  High voltage of 500 V is applied to the HV lines on the hybrid. The currents drawn by, and the temperatures of the hybrids are monitored every few minutes.  Every few hours a test to establish correct functionality of the hybrid is performed, so that if problems do develop the time structure of the failures can be observed.

### Final Electrical Test of Hybrids

After the long term test, the full electrical tests (Characterisation Sequence – see sect. ???) are performed. Results are stored in the database

## QA during module assembly

In order to ensure the quality of the module, important conditions must be met during the assembly. These conditions and rules are described in another section of this document (Assembly Criteria). Here, the tests to verify the module quality are listed.

After detectors are glued to the spine, IV curve of each detector is measured at 20 deg C up to 400 V.  Leakage current at 350 V should not exceed specification current.

Furthermore, XY metrology survey is performed at 20 deg C.

## QA of the completed module

Fully completed (glued and bonded) module is subject to thorough testing. The following QA steps are performed on fully assembled modules.  The criteria for accepting a module as ‘good’ are for initial use, and will evolve with experience.

### IV curve

IV curve of the whole 4-detector assembly via hybrid connector is measured, similarly to Sect. 4.3. Only modules with I(leak) < 4 x I(specs) are considered as ‘good’.

### Thermal cycling

Every module will be subject to thermal cycling between –30 and 50 deg C (temperatures referring to hybrid thermistor), without applying bias, or power. The ramp up/down times will be approximately 30 minutes (2-3oC/minute) and the soak time about 30 minutes at each temperature.  The total test time will therefore be about 20 hours.

### Metrology survey

After thermal cycling, final metrology survey at 20 deg C is done – both XY(in-plane)  and Z (out-of-plane).

The metrology procedure is described in a separate document.

### IV curve of the completed module

After thermal cycling and  final metrology survey, the I-V curve is recorded again. Only modules with I(leak) < 4 x I(specs) are considered as ‘good’.

### Long-term leakage current stability

All modules will be tested for long-term leakage current stability over a 24 hour period in an environmental chamber containing cold dry air (nitrogen).  The ASICs are powered, clocked and triggered, and the detector bias is maintained at 150 V, with the current monitored every 15 minutes over the period.  The temperature is adjusted so that it is 10oC measured by the hybrid thermistors. The maximum increase in leakage current over the period should be less than 4 A for a good module, after an initial settling time of 5 minutes.

This test can be performed in parallel with the long-term electrical test on modules (see below).

### Module electrical confirmation test

In this test basic electrical functionality is tested using the Confirmation Sequence.

### Long-term electrical stability

In addition to detailed electrical testing and characterisation, a longer duration test will be performed on assembled modules at the ATLAS operating temperature.  This test verifies that modules will function electrically at reduced temperatures.

The test consists of an extended (24 hour) run at reduced temperature, defined to be –10oC as measured by the hybrid thermistors.  During the test the hybrids are clocked and triggered.  The currents drawn by, and the temperatures of, the hybrids are monitored continuously.

??? Every few hours a Confirmation Sequence is performed.

This test may be performed in parallel with the long-term leakage test.

### Final electrical characterization tests

After the long term test, the full electrical tests of each module (Characterisation Sequence – see below) are performed. Results are stored in the database.

## Sampling QA on completed modules

The fraction of modules to be used initially for the sub-sample tests is indicated in each of the following sections.

### Laser Tests

A subset of the produced modules will also be submitted to scanning Laser tests.  A focused LASER of wavelength e.g. 1050 nm is mounted on a x-y stage and scanned over the module.  This test provides very precise position information, hence the correct functionality of all channels can be verified. Furthermore, bonding errors can be revealed by this test. It is anticipated, that at least one module per production site will be scanned by laser.

### Irradiation Tests

A very small sample of the completed modules (approximately 10 per annum during the construction period) will be fully and uniformly irradiated in the SCT facility at the CERN PS to a fluence of 3x1014 pcm-2 24 GeV/c protons.  They will be annealed for 7 days at 25oC following the irradiation and then checked for mechanical integrity and for noise performance, for full ASIC functionality and for detector leakage current when run cold at the SCT operating temperature.  Several of these modules will also be tested in the beam for signal:noise and efficiency performance.

### Beam Tests

A small number (approximately 10 per annum during the construction period) of the end cap modules will be fully tested in the H8 beam at CERN, with a magnetic field and with varying angles of incidence of the particles to check their continued performance characteristics.

### Source Tests

It is anticipated that a fraction of the modules will, at least initially, be read out in the laboratory when exposed to a beta source.  This will allow signal:noise values to be confirmed for the different batches of delivered ASICs.

## Definition of Electrical Hybrid and Module Tests

An extensive suite of both hardware and software has been developed to facilitate module testing.  The readout system is based around the VME modules CLOAC, MuSTARD and SLOG.  The prototype low voltage and high voltage modules, SCTLV and SCTHV, are also part of the system.

Two largely automated series of tests have been devised to simplify the testing procedure, as outlined in table 1.  The Characterisation Sequence aims to perform the full characterisation of a hybrid or module whereas the shorter Confirmation Sequence provides a reduced set of information.  The Confirmation Sequence ensures that the digital part of the ASICs is functioning, none of the critical wire-bonds have been damaged and that the basic analogue performance of a module has not deteriorated.  It is anticipated that the Confirmation Sequence would be repeated at regular intervals during the long-term tests and each time that a hybrid or module is shipped between institutes.

|  | Characterisation | Confirmation |
| --- | --- | --- |
| Power On Tests / Verification of Response to Hard Reset |  |  |
| Clock and Command Reception Test |  |  |
| Bypass Functionality Test |  |  |
| Pipeline Efficiency Test |  |  |
| Strobe Delay Scan |  |  |
| Three Point Estimation of Gain, Noise and Offset |  |  |
| TrimRange Scan |  |  |
| Determination of the Response Curve |  |  |
| Noise Occupancy Scan |  |  |
| Timewalk Scan |  |  |

Table 1: The Characterisation and Confirmation Sequences

In general, the analogue performance of a module/hybrid is measured with respect to the internal calibration circuitry of the ABCD3T chip.  Hence it is necessary to make a correction for the variation of the calibration capacitance between batches of ASIC wafers.

There follows a brief description of the individual tests to be performed as part of the electrical QA procedure. The description contains the method in general terms, the purpose of the tests and in quantitative terms the criteria for PASS/FAIL cuts.  A summary of the results of each test will be recorded in the SCT database.

A module is classed as `good’ if at least 99% of its readout strips will operate efficiently with low noise occupancy at 1fC threshold.

All parametrical tests (response curve, noise occupancy, etc.) should be performed using  the primary clock&control readout channel, i.e. via DORIC chip. Them modules should be read out via VDC chip. This will be done using the optical readout or electrical emulation.

### Power on Tests / Verification of Response to Hard Reset

The module/hybrid is clocked and the power is switched on.  The operator must verify that each data-link responds with CLK/2 and that, after the chips have been configured, the clock feed-through signal stops.  The analogue and digital currents are then recorded.  Finally Hard Reset is issued to bring back the CLK/2 signal.

This test verifies that the Clock, Command and Hard Reset signals are received correctly, that the chips can be configured and that the current consumption is reasonable.  The test will identify modules/hybrids with severe failures.  Every module must pass this test without error.

This is the only test that would normally require operator intervention.

### Clock and Command Reception / Addressing Error Test

The chips are configured to return the contents of the Mask Register and a burst of triggers is issued for each of the Primary and Redundant Clock and Command options.  Prior to each event, a different bit pattern is loaded in the Mask Register such that consecutive events are not the same.

By comparing the received data with expectation it is verified that both the Primary and Redundant Clock and Command signals are received correctly and that the top address bit of each chip changes as the Clock/Command source is varied, as specified in the module design.  This test will identify modules/hybrids with faulty command reception or addressing errors.  Modules/hybrids with such defects would be considered to have failed pending further investigation and possible rework.

### Bypass Functionality Test

A trigger burst is recorded with the module/hybrid programmed to each of a number of different configurations, sufficient to exercise all data/token passing links between the chips.  In each case the chips are configured to return the contents of the Mask Register such that the expected data is accurately known.  The test is repeated across a range of digital supply voltages.

This test determines the minimum value of the digital supply voltage needed for each of the data/token passing links to work.  Any link that did not work at the designated supply voltage, and which could not be identified as being due to a missing wirebond and subsequently repaired, would cause a module/hybrid to be rejected.

### Pipeline Efficiency Test

For this test, a Soft Reset command is sent to reset the pipeline followed a certain number of clock periods later by a Pulse Input Register command and L1A trigger.  In this way, a known pattern is injected into a given location in the pipeline.  By varying the distance between the Soft Reset and Pulse Input Register commands it can be verified that each of the eleven blocks within the pipeline is free of defects.

Zero occupancy for a particular number of clock periods between the Soft Reset and Pulse Input Register commands would indicate a dead cell in the corresponding block of the pipeline.  Zero occupancy for all values would indicate a dead channel.  Modules/hybrids with a large number of dead Pipeline cells or dead channels will be rejected.

### Strobe Delay Scan

This scan is performed to determine the correct Strobe Delay setting, corresponding to the timing of the charge injection pulse, to be used during the Analogue Tests.

### Three Point Estimation of Gain, Noise and Offset

Threshold scans are taken for three injected charges to facilitate a quick measurement of gain, noise and the discriminator offset.  Pathological channels are categorised as FAULTY if the defect would result in the channel having a reduced but non-zero detection efficiency in ATLAS, or as LOST if the defect would result in the channel having zero efficiency:

- Lost: Dead, Stuck, Unbonded or Noisy channels

- Faulty: Inefficient, Low Gain or Partially Bonded channels

Modules/hybrids having any chips with abnormal gain or high noise will be rejected, for potential re-work, as will those with large numbers of pathological channels.

### Trim Range Scan

For each of the four possible TrimRange settings, a series of Threshold scans are performed for a subset of the sixteen possible TrimDAC settings, all with 1fC injected charge.  For each TrimRange setting a straight line is fitted to the data for each channel to characterise the TrimDAC response and to determine the TrimDAC slope.  The number of trimmable channels and the spread of the resultant trimmed thresholds are also recorded.  The optimised TrimDAC settings and a list of channels to be masked are produced for use in the subsequent analogue tests.

The chips used to build modules will have been selected such that all channels may be trimmed using the smallest TrimRange.  Modules, which do not meet this specification on at least 99% of channels, will be rejected, for potential rework, as will those where a particular TrimRange has a slope other than that expected.

### Response Curve

Threshold scans are performed for a series of input charges and, for each channel, an appropriate function is fitted to the resulting response curve.  From this the Gain, Noise and discriminator Offset are extracted.

The parameters from the fit are stored since they describe the correspondence between the Threshold, in mV, and input charge, in fC.  The categorisation of pathological channels is repeated as described for the Three Point Gain.  Modules/hybrids with a large number of pathological channels will be rejected.

### Noise Occupancy Scan

A high statistics Threshold scan is performed at the nominal ATLAS trigger rate of 100kHz, without any injected charge, to determine the Noise Occupancy of each channel as a function of Threshold.  The analogue and digital current consumption as a function of Threshold is recorded.

Channels with high Noise Occupancy will be added to the list of masked channels.

### Timewalk Scan

This test performs a series of Strobe Delay scans with the Threshold set to 1 fC, varying the input charge from 1.25 to 10 fC. In each case a fit is made to the rising edge of the pulse to determine the Strobe Delay value needed to obtain 50% occupancy.

The Timewalk is defined as the time variation in the crossing of a threshold of 1fC over a signal range of 1.25 to 10.0fC.  This parameter is calculated and recorded.

## QA Organisation

Here the list of QA sites with their individual tasks will be given

### Site qualification

To be done

### Infrastructure and equipment

To be done

# Project Schedule

## End-Cap Construction Schedule

## Module Production Responsibilities

## Component Supplies

## Timetable

# Prototype Module Results

## Pre-irradiation at Individual Module Level

### Mechanical and Thermal

#### Metrology

#### Thermal Cycling Results

#### Cooling Tests

### Electrical

#### Stability

#### Noise Occupancy

#### S/N

## Pre-irradiation at Combined Modules Level

### System Test

## Post-irradiation

### Mechanical Integrity

### Electrical Performance

# References

This description is currently available in the draft of Didier Ferrere, available at http://dpnc.unige.ch/atlas/ferrere/assembly/documents/FDRdraft.pdf

The draft of a metrology document written by Steve Snow is available at http://www-ucjf.troja.mff.cuni.cz/~sct/QA/docs/Stevesurv2.pdf

http://sct.home.cern.ch/sct/sctdaq.html

The tests are described fully in http://hepwww.rl.ac.uk/atlas-sct/documents/Electrical_Tests.htm