21.01.2002

# DCS Sensor Set-up

## Cable connection and pin-out

The figure below shows the pin-layout of the sensors on the two green boards in the DCS prototype connected by the thick orange cable (original drawing resides in dcs_pins.cdr).

## LMB pin-out

The channels and assignment of the two 34-pin connectors of the LMB boxes is as follows:

LMB-A, 1st connector:

| Channels | Pins | Sensor |
| --- | --- | --- |
| 0 | 1,2 | 100 (R1) |
| 1 | 3,4 | PT100#6 (T6) |
| 2 | 5,6 | PT100#1 (T1) |
| 3 | 7.8 | PT100#2 (T2) |
| 4 | 9,10 | PT100#heater (Thtr) |
| 5 | 11,12 | Gas dew point 2 |
| 8 | 17,18 | Gas dew point 1 |

LMB-A, 2nd connector is currently not used.

LMB-R, 1st connector:

| Channels | Pins | Sensor |
| --- | --- | --- |
| 1 | 1-4 | Humidity1, R1 (H1.1) |
| 3 | 5-8 | Humidity1, R2 (H1.2) |
| 5 | 9-12 | Humidity2, R1 (H2.1) |
| 7 | 13-16 | Humidity2, R2 (H2.2) |
| 9 | 17-20 | PT1000#11 (T11) |
| 11 | 21-24 | PT1000#12 (T12) |
| 13 | 25-28 | PT1000#13 (T13) |
| 15 | 29-32 | PT1000#14 (T14) |

LMB-R, 2nd connector:

| Channels | PinsError: Reference source not found | Sensor |
| --- | --- | --- |
| 17 | 1-4 | PT1000#15 (T15) |
| 19 | 5-8 | PT1000#16 (T16) |
| 21 | 9-12 | PT1000#7 (T7) |
| 23 | 13-16 | PT1000#8 (T8) |
| 25 | 17-20 | PT1000#9 (T9) |
| 27 | 21-24 | PT1000#10 (T10) |
| 29 | 25-28 | PT1000#17 (T17) |
| 31 | 29-32 | PT1000#18 (T18) |

## Currently used sensors

The following sensors are currently used in the SCT DCS system:

| Hardware name | LabVIEW name | System | Function |
| --- | --- | --- | --- |
| R1 | ### not shown | DCS internal use | PT100’s current control |
| PT100#6 | B-Amb#left | Barrel sector | Ambient temperature |
| PT1000#17 | B-Amb#centre | Barrel sector | Ambient temperature |
| PT1000#18 | B-Amb#right | Barrel sector | Ambient temperature |
| PT100#1 | B-Amb#back | Barrel sector | Ambient temperature |
| PT100#2 | B-Amb#front | Barrel sector | Ambient temperature |
| PT100#heater | Heater (not shown) | Cryophasing equipment | Heater internal temperature |
| Gas dew point 1 | B-N2-Dew | Barrel sector | Dewpoint of incoming N2 |
| Gas dew point 2 | F-N2-Dew | Forward sector | Dewpoint of incoming N2 |
| PT1000#11 | B-PipeA-in | Barrel sector | Cooling pipe temperature |
| PT1000#12 | B-PipeB-in | Barrel sector | Cooling pipe temperature |
| PT1000#13 | B-PipeC-in | Barrel sector | Cooling pipe temperature |
| PT1000#14 | B-PipeA-out | Barrel sector | Cooling pipe temperature |
| PT1000#15 | B-PipeB-out | Barrel sector | Cooling pipe temperature |
| PT1000#16 | B-PipeC-out | Barrel sector | Cooling pipe temperature |
| PT1000#7 | F-Amb-back | Forward sector | Ambient temperature |
| PT1000#8 | F-Amb-front | Forward sector | Ambient temperature |
| PT1000#9 | F-Pipe-in | Forward sector | Cooling pipe temperature |
| PT1000#10 | F-Pipe-out | Forward sector | Cooling pipe temperature |
| Humidity1 | BarrelHum | Barrel sector | Ambient humidity |
| Humidity2 | ForwdHum | Forward sector | Ambient humidity |
| ### no hardware | Dewpoint 1 | Barrel sector | Dewpoint from BarrelHum |
| ### no hardware | Dewpoint 2 | Forward sector | Dewpoint from ForwdHum |

 The first two pins in a group of four LMB-R pins are used for current in-/output, the last two for voltage measurement.