Lecroy 8901A Gpib Interface

________________________________________________________________________________

Introduction:

This instrument driver provides programming support for the LeCroy 8901A GPIB

Crate Controller.

________________________________________________________________________________

Assumptions:

To successfully use this module, the following conditions must be met:

-  the instrument is connected to the GPIB.

-  the GPIB address supplied to the initialize function must match the

GPIB address of the instrument.

________________________________________________________________________________

Error and Status Information:

Each function in this instrument driver returns a status code that

either indicates success or describes an error or warning condition.

Your program should examine the status code from each call to an

instrument driver function to determine if an error occurred.

The general meaning of the status code is as follows:

Value                  Meaning

-------------------------------

0                      Success

Positive Values        Warnings

Negative Values        Errors

The description of each instrument driver function lists possible error

codes and their meanings

________________________________________________________________________________

How To Use This Document:

Use this document as a programming reference manual.

It describes each function in the

Lecroy 8901A Gpib Interface

instrument.  The functions appear in alphabetical order, with a description

of the function and its C syntax, a description of each parameter, and a list

of possible error codes.

________________________________________________________________________________

Function Tree Layout:

Lecroy 8901A Gpib Interface            Function Name:

Initialize                                  lc8901a_init

Mode                                        lc8901a_Mode

Write                                       lc8901a_Write

Read                                        lc8901a_Read

Initialize Clear Inhibit                    lc8901a_InitClearInhibit

Set SRQ                                     lc8901a_SetSRQ

Wait For SRQ                                lc8901a_WaitSRQ

Error Message                               lc8901a_errorMessage

Close                                       lc8901a_close

________________________________________________________________________________

Lecroy 8901A Gpib Interface

This instrument module provides programming support for the Lecroy

8901A.  The module is divided into the following functions:

________________________________________________________________________________

The following functions are in alphabetical order.

________________________________________________________________________________

lc8901a_close

ViStatus lc8901a_close (ViSession instrumentHandle);

Purpose

This function performs the following operations:

viClose (instrSession) and viClose (rmSession).

Notes:

(1) The instrument must be reinitialized to use it again.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  None

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

lc8901a_errorMessage

ViStatus lc8901a_errorMessage (ViSession instrumentHandle, ViStatus statusCode,

ViChar message[]);

Purpose

This function takes the Status Code returned by the instrument driver

functions, interprets it and returns it as a user readable string.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  VI_NULL

statusCode

Variable Type       ViStatus

This control accepts the Status Code returned from the instrument

driver functions.

Default Value:

0 - VI_SUCCESS

message

Variable Type       ViChar[]

This control returns the interpreted Status Code as a user readable

message string.

Notes:

(1) The array must contain at least 256 elements ViChar[256].

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

3FFF0085  VI_WARN_UNKNOWN_STATUS

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

lc8901a_init

ViStatus lc8901a_init (ViRsrc resourceName, ViInt16 transferMode,

ViInt16 dataWord, ViSession *instrumentHandle);

Purpose

This function performs the following initialization actions:

- Opens a session to the Default Resource Manager resource and a session

to the specified device using the interface and address specified in the

Resource_Name control.

- Sends initialization commands to the instrument that set the mode for

subsequent writes and reads.

- Returns an Instrument Handle which is used to differentiate between

different sessions of this instrument driver.

- Each time this function is invoked a Unique Session is opened.  It is

possible to have more than one session open for the same resource.

Parameter List

resourceName

Variable Type       ViRsrc

This control specifies the interface and address of the device that

is to be initialized (Instrument Descriptor). The exact grammar to be

used in this control is shown in the note below.

Default Value:  "GPIB::1"

Notes:

(1) Based on the Instrument Descriptor, this operation establishes a

communication session with a device.  The grammar for the Instrument

Descriptor is shown below.  Optional parameters are shown in square

brackets ([]).

Interface   Grammar

------------------------------------------------------

GPIB        GPIB[board]::primary address[::secondary address]

[::INSTR]

The GPIB keyword is used with GPIB instruments.

The default value for optional parameters are shown below.

Optional Parameter          Default Value

-----------------------------------------

board                       0

secondary address           none - 31

transferMode

Variable Type       ViInt16

This control specifies the transfer mode for all subsequent read and

write operations for the LC 8901A.

Valid Range:

Single Transfer        0 (Default)

Block Read             3

High Speed Block Read  6

dataWord

Variable Type       ViInt16

This control specifies the length of data words to be used for

subsequent writes and reads for the LC 8901A.

Valid Range:

8 Bit     0

16 Bit    1  (Default)

24 Bit    2

instrumentHandle

Variable Type       ViSession (passed by reference)

This control returns an Instrument Handle that is used in all

subsequent function calls to differentiate between different sessions

of this instrument driver.

Notes:

(1) Each time this function is invoked a Unique Session is opened.

It is possible to have more than one session open for the same

resource.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0803  Invalid Response VI_ERROR_INSTR_INTERPRETING_RESPONSE

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

lc8901a_InitClearInhibit

ViStatus lc8901a_InitClearInhibit (ViSession instrumentHandle, ViInt16 command);

Purpose

This function causes the LC 8901A to assert a CAMAC Initialize (Z), Clear

(C) or Inhibit (I). In addition Inhibit can be deasserted through with

this function.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  VI_NULL

command

Variable Type       ViInt16

This variable selects the CAMAC command or condition that is to be

asserted.

Valid Range:

Initialize (Z)       0  (Default)

Clear (C)            1

Initialize and Clear 2

Inhibit (I)          3

De-assert Inhibit    4

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 out of range.

BFFC0803  Invalid Response VI_ERROR_INSTR_INTERPRETING_RESPONSE

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

lc8901a_Mode

ViStatus lc8901a_Mode (ViSession instrumentHandle, ViInt16 transferMode,

ViInt16 dataWord);

Purpose

This function is used to specifiy the mode for all subsequent writes and

reads to and from the LC 8901A. The transfer mode and length of data are

configured through this function.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  VI_NULL

transferMode

Variable Type       ViInt16

This control specifies the transfer mode for all subsequent read and

write operations for the LC 8901A.

Valid Range:

Single Transfer        0 (Default)

Block Read             1

High Speed Block Read  2

dataWord

Variable Type       ViInt16

This control specifies the length of data words to be used for

subsequent writes and reads for the LC 8901A.

Valid Range:

8 Bit     0

16 Bit    1  (Default)

24 Bit    2

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 out of range.

BFFC0003  Parameter 3 out of range.

BFFC0803  Invalid Response VI_ERROR_INSTR_INTERPRETING_RESPONSE

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

lc8901a_Read

ViStatus lc8901a_Read (ViSession instrumentHandle, ViInt16 f, ViInt16 a,

ViInt16 n, ViInt32 blockCount, ViInt32 timeoutValuems,

ViInt32 data[], ViBoolean *x, ViBoolean *q);

Purpose

This function is used to read data from the Camac Crate. The module to

read from and the read operation to be performed are specified by the F,

A, and N parameters.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  VI_NULL

f

Variable Type       ViInt16

This control specifies the function code (F) that is to be written.

Default Value:  0

Valid Range:  0 to 7

a

Variable Type       ViInt16

This control specifies the Subaddress (A).

Default Value:  0

Valid Range:  0 to 15

n

Variable Type       ViInt16

This control specifies the Station Number (N) or Crate Slot Number

that is intended to receaive the command.

Default Value:  1

Valid Range:  0 to 23

blockCount

Variable Type       ViInt32

This variable specifies the maximum number of blocks expected by the

user. If the exact number is not known a high value should be used to

ensure that all the blocks are read back. The "Timeout Value" should

be adjusted accordingly if a large "Block Count" in entered.

Default Value:  100

Note:

This varable will be ignored, but must still be declared, if the LC

8901A is in single transfer mode.

timeoutValuems

Variable Type       ViInt32

This variable specifies the time out value for the read. This

specifies the maximum time in which the number of blocks specified by

the "Block Count" variable should be read. If the blocks are read

under this time, the read will terminate automatically. If the blocks

aren't read in the specified time a time out error will be returned

by the "error out" cluster.

Default Value:  3000 ms

Note:

This variable will be ignored if the LC 8901A is in single transfer

mode (the variable must be declared).

data

Variable Type       ViInt32[]

This array returns the read data as signed 32-bit integers. Each

element of the array represents one piece of data, the first data

being in the first array element (index =0). If the instrument is in

single transfer mode the first element of the array will contain the

data value read.

x

Variable Type       ViBoolean (passed by reference)

This variable indicates whether an X condition exists. A return valud

of 1 indicates that an X conditino exists.

q

Variable Type       ViBoolean (passed by reference)

This variable indicates whether an Q condition exists. A return value

of 1 indicates that a Q condition exists.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 out of range.

BFFC0003  Parameter 3 out of range.

BFFC0004  Parameter 4 out of range.

BFFC0005  Parameter 5 out of range.

BFFC0006  Parameter 6 out of range.

BFFC0803  Invalid Response VI_ERROR_INSTR_INTERPRETING_RESPONSE

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

lc8901a_SetSRQ

ViStatus lc8901a_SetSRQ (ViSession instrumentHandle, int SRQCondition);

Purpose

This function is used to configure SRQ conditions for the LC 8901A.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  VI_NULL

SRQCondition

Variable Type       int

This variable specifies the SRQ Condition that is to be set for the

LC 8901A.

Valid Range:

Disable SRQ                         0  (Default)

Enable SRQ on LAM                   1

Enable SRQ on Q=0                   2

Enable SRQ on LAM or Q=0            3

Enable SRQ on X=0                   4

Enable SRQ on LAM or X=0            5

Enable SRQ on Q=0 or X=0            6

Enable SRQ on LAM or Q=0 or X=0     7

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 out of range.

BFFC0803  Invalid Response VI_ERROR_INSTR_INTERPRETING_RESPONSE

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

lc8901a_WaitSRQ

ViStatus lc8901a_WaitSRQ (ViSession instrumentHandle, long timeOut,

short lamResponse[], ViBoolean *x, ViBoolean *q);

Purpose

This function is used to wait for the SRQ condition enabled by the "Set

SRQ" function. The time to wait for the SRQ condition can be specified.

After the SRQ condition has been detected the wait will terminate, SRQ

will be disabled and the Lam state of each module will be returned.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  VI_NULL

timeOut

Variable Type       long

This paramater is used to set the timeout to wait for a Service

Request (SRQ) in ms.

Default Value:  30000 ms

Valid Range:  Select the largest value that you wish to wait for. A

time out will not occur until this time has passed and very large

values may result in unnecessarily long wait times in case of errors.

lamResponse

Variable Type       short []

This short integer array is used to return the LAM Response for the

SRQ. The first element of the array (index 0) corresponds to slot 1,

the second to slot 2 and so on. If the element of any array is 1 a

LAM was issued by the associated slot. A 32 element array must be

defined and passed through the function.

Note:

1) The array must contain 32 elements short[32].

x

Variable Type       ViBoolean (passed by reference)

This variable indicates whether an X condition exists. A return value

of 1 indicates that an X condition exists.

q

Variable Type       ViBoolean (passed by reference)

This variable indicates whether an Q condition exists. A return value

of 1 indicates that a Q condition exists.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 out of range.

BFFC0803  Invalid Response VI_ERROR_INSTR_INTERPRETING_RESPONSE

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.

________________________________________________________________________________

lc8901a_Write

ViStatus lc8901a_Write (ViSession instrumentHandle, ViInt16 f, ViInt16 a,

ViInt16 n, long data);

Purpose

This function is used to write to the LC 8901A. The "Data" variable will

be ignored for functions that do not require it. The function should not

be used to send the command to perform a read from the Crate Controller,

the command and read are performed together in the "lc8901a_Read"

function.

Parameter List

instrumentHandle

Variable Type       ViSession

This control accepts the Instrument Handle returned by the Initialize

function to select the desired instrument driver session.

Default Value:  VI_NULL

f

Variable Type       ViInt16

This control specifies the function code (F) that is to be written.

Default Value:  16

Valid Range:  8 to 31

a

Variable Type       ViInt16

This control specifies the Subaddress (A).

Default Value:  0

Valid Range:  0 to 15

n

Variable Type       ViInt16

This control specifies the Station Number (N) or Crate Slot Number

that is intended to receaive the command.

Default Value:  1

Valid Range:  0 to 23

data

Variable Type       long

This integer specifies the data that is to be sent with the specified

command (FAN). This integer is passed as a signed 32 bit number, the

function will automatically convert the data to the appropriate

format, selected by the "lc8901a_Mode" function.

Return Value

This control contains the status code returned by the function call.

Status Codes:

Status    Description

-------------------------------------------------

0  No error (the call was successful).

3FFF0005  The specified termination character was read.

3FFF0006  The specified number of bytes was read.

BFFC0002  Parameter 2 out of range.

BFFC0003  Parameter 3 out of range.

BFFC0004  Parameter 4 out of range.

BFFC0803  Invalid Response VI_ERROR_INSTR_INTERPRETING_RESPONSE

BFFF0000  Miscellaneous or system error occurred.

BFFF000E  Invalid session handle.

BFFF0015  Timeout occurred before operation could complete.

BFFF0034  Violation of raw write protocol occurred.

BFFF0035  Violation of raw read protocol occurred.

BFFF0036  Device reported an output protocol error.

BFFF0037  Device reported an input protocol error.

BFFF0038  Bus error occurred during transfer.

BFFF003A  Invalid setup (attributes are not consistent).

BFFF005F  No listeners condition was detected.

BFFF0060  This interface is not the controller in charge.

BFFF0067  Operation is not supported on this session.