30.10.2001

# Cryophasing Manual

CONTENTS

1. Principle of Operation

2. LabVIEW Application

3. What to Do When

4. References

## Principle of Operation

### 1.1 Hardware

The following hardware components are necessary to perform the cryophasing:

- Huber chiller controlled from the DCS PC serial port via the SUB-D connector on the front panel. The control allows to change the set-point and read the temperature of the cooling liquid and the chiller status remotely [3.1].

- Cooling circuit in the freezer (“coil”) interfaced to the chiller and the cooling pipe(s) on the sector via two valve-panels in front of the freezer.

- Freezer controlled from the DCS PC parallel port (LPT1) via a relay (freezer on/off) [3.2].

- Heating element in the cooling circuit connected to the sector pipes inside the freezer; driven with a 24V/4A power supply and controlled via relays from the DCS PC parallel port (heater on/off and full/half power) [3.2].

- The DCS system [3.3] needed for temperature and humidity monitoring.

- Dry nitrogen supply (manual control) with moisture measurement.

Before starting operation make sure that all these components are connected and working properly.

### 1.2 Procedure

The cool-down and warm-up procedures are performed following the steps listed below.  Figure 1 shows the DCS reading during cryophasing.

Figure 1: DCS readings during cryophasing. The numbers indicate the steps of cool-down and warm-up as listed in section 1.2.

The process of cool-down should be performed in the following steps:

- Put the sector into the freezer at room temperature and close and (optionally) seal the cover.

- Start DCS monitoring [3.3].

- Start dry gas flow, ie flush with dry nitrogen.

- When a humidity-plateau is reached start cooling the coil (brass tubes inside the freezer).

- If necessary use freezer, switched on in short intervals only.

- When the dew point is well below desired temperature turn on freezer. Set chiller to room temperature and wait till the cooling liquid temperature is near the freezer ambient temperature.

- Cool sector pipes. Stop coolant flow through the coil. Cool down cooling liquid again.

- When desired temperatures are stable start measurements. Maintain ambient temperature via the freezer control.

The warm-up follows these steps:

- Turn off freezer and set chiller to room temperature.

- Turn on heater at half power.

- When cooling pipes on sector reach room temperature carefully open valves to the coil and continue heating with full power.

- When ambient temperature inside freezer is well above the dew point outside, the freezer cover can be opened.

## LabVIEW Application

### 2.1 Start

Pressing the “Open Cryophase Panel” button in the DCS main panel opens the application window. This program automatically performs most of above steps. Some operations have to be done manually, the program will pop up instructions accordingly.

The “selected system”-control at the top-right of the panel selects the humidity and dew point values used for process control during cryophasing and the user can select barrel or forward for the respective systems. This control must be set before operation starts. After some initialisations (according message in the bottom field) two buttons to the left appear. They are used to start the cool-down procedure or stop the application, respectively (see below). A warm-up and an abort button for respective actions during operation replace them. In case of errors  like DCS system not running properly  a display above these buttons will show according messages. To the right there are two controls for the desired pipe and ambient temperature after cool-down and two LED’s indicating the status of freezer and heater. In the centre of the panel you find two thermometers indicating the current temperature of the cooling liquid (chiller) and the ambient temperature in the freezer with their respective digital displays.

### 2.2 Cool-down

After start-up of the LabVIEW application (see above) set the ambient and pipe temperature controls according to your needs, press the “cool-down” button, and wait for instructions. Manual operations like starting the dry gas flow or opening valves are ordered by message windows which appear whenever such an action is needed.

The program follows the items listed in section 1.2. When a stable condition at the desired temperatures is reached a “ready for data taking” message appears in the bottom field and temperatures are automatically maintained at the current level. At the same time the “warm-up” button replaces the “cool-down” button and the “abort” button disappear. Further cool-down runs can only be started after successive warm-ups.

### 2.3 Warm-up

This phase can only be entered after a cool-down was run successfully. Press the “warm-up” button and wait for instructions. As for the cool-down, most operations are done automatically, but some valves have to be opened/closed by hand. After successful warm-up, the message “freezer can be opened now” appears in the bottom field and the “cool-down” and “stop” buttons re-appear.

### 2.4 Stop

The option to stop the application, ie to perform a controlled termination, is only allowed after a successful warm-up or an abort or straight after initialisation. A “stop” button should appear on the panel, pressing which the termination is run and after several seconds the application stops automatically. The panel can be closed from the Windows-menu or by pressing Ctrl-W.

### 2.5 Error Handling

In case of errors – no DCS data, DCS alarms or fatal errors, or chiller errors – both cool-down and warm-up procedures are interrupted automatically. The operation of the heater will be disabled to avoid overheating and both freezer and chiller operation will be disabled with the chiller being set to room temperature. A “procedure interrupted” message appears in the bottom field and the cause of the error, if it is not a DCS sensor alarm, is displayed, too. Special caution should be taken on how to continue since the freezer might be below the room dew point already in which case there is a risk of condensation on the sector.

Any process can be interrupted during operation with the “abort” button which should bring the program into a state that allows to re-start the cool-down procedure or to stop the program entirely and has an effect like an interruption triggered by DCS/chiller errors.

## What to Do When

### 3.1 Errors from the Chiller

- If  the message “Error” shows up in the chiller display the level in the cooling bath is probably too low; fill with cooling liquid and restart. As well the colling water flow might be too low or even off. Check the valves on the water suppliy pipes and the valve in the back of the chiller (see next item). If this does not help check the manual [3.1] for other error sources.

- If  the message “Cond” is displayed the water flow through the condensor is too low. Open bottom panel on the back of the chiller and tighten top screw on the input valve (brass set-up in the middle).

- The internal coolant temperature is very different (more than a few degrees) from the set-point: check display for “Cond” message and follow above instructions. If this does not help the coolant is perhaps frozen – abort cooling and wait till coolant liquifies again.

### 3.2 Errors from DCS

- Sensor errors: most likely due to mistakes during operation or faults in the equipment/bugs in the software; try to fix the problem causing the error flag (fix broken sensor, take unreliable sensor off the limit list, etc), clear DCS error flag or re-start DCS, and re-start cryophasing.

- DCS failure: see [3.3] to fix DCS problems; re-start cryophasing after DCS is back running.

## References

- Huber Intelligent Chiller Control, Operating Instructions; Huber Chiller Remote Interface Manual: Chiller.doc

- Freezer and Heater Interface Manual: Parallel.doc

- SCT DCS Prototype - LabVIEW Application User Manual: Sct_Dcs_LabVIEW.doc

Comments to J.Grosse-Knetter (joern.grosse-knetter@cern.ch)